## Accuracy & CMRR
Calibration procedure Linear Audio AutoRanger v1.5
Part 1 – Display accuracy
Step 1a | Switch AR off, then on while pressing the BAL pushbutton. AR starts up at 1V nominal setting
Connect AC DMM to TP7 and TP5 (gnd)
Switch AR off, then on while keeping SE input push-button pressed. AR will go into cal mode.
Insert 1kHz 1V signal at SE input as measured on DMM at TP7, as accurate as possible
Press HOLD button to complete cal procedure.
Remove 1kHz 1V signal
Step 1b | Switch AR off, then on while pressing the HOLD pushbutton. AR starts up at 0.4V nominal setting
Connect AC DMM to TP7 and TP5 (gnd)
Switch AR off, then on while keeping SE input push-button pressed. AR will go into cal mode.
Insert 1kHz 0.4V signal at SE input as measured on DMM at TP7, as accurate as possible
Press HOLD button to complete cal procedure.
Remove 1kHz 0.4V signal
Part 2 – Common mode suppression
Step 1 | Switch AR off, then on while pressing the BAL pushbutton. AR starts up at 1V nominal setting
Connect AC DMM to TP7 and TP5 (GND)
Connect TP1 to TP9
Set AR to SE input
Input ~1V signal @ 1kHz to the SE input, as measured at TP7, observe AR settling at 0dB
Set AR to BAL input mode.
Connect AC DMM to TP7
Adjust RV1 for minimum indication on the DMM.
Step 2 | Set AR to HOLD, then repeatedly press SE input to switch to -48dB setting
Adjust signal generator for maximum output but less than 100VRMS at TP9
Adjust RV2 for minimum indication on the DMM as measured on TP7.
Remove input signal
Remove short between TP1 and TP9

## SE flatness
Calibration procedure Linear Audio AutoRanger v1.5 |  | Part 3 – SE mode frequency response
Warning: DO NOT change the contents of the green highlighted cells!
The black numbers in the J column have to be replaced with what you measure!
Step 1 -48dB cal | Set AR-1 to SE mode
Connect DMM gnd probe to TP5
Set AR to HOLD then press SE input until unit switches to -48dB setting
Input ~10V @ 10kHz at SE input as measured at TP9
| Enter measured level at TP9 in the 10kHz Vin -48dB SE cell below |  | 10.58
| Connect AC DMM to TP7, enter measured level at 10kHz Vout in -48dB SE cell below |  | 0.042
| Switch generator to 100kHz, enter value measured at TP9 |  | 10.583
| Connect AC DMM to TP7, Adjust VC4 for DMM indication of green highlighted value |  | 0.042012
Step 2 -32dB cal | With HOLD still activated, press SE input until unit switches to -32dB setting
Input ~10V @ 10kHz at SE input as measured at TP9
| Enter measured level in the 10kHz Vin -32dB SE cell below |  | 10.582
| Connect AC DMM to TP7, enter measured level at 10kHz Vout in -32dB SE cell below |  | 0.2655
| Switch generator to 100kHz, Enter value measured at TP9: |  | 10.58
| Connect AC DMM to TP7, adjust VC6 for DMM indication of green highlighted value |  | 0.26545
Step 2 -16dB cal | With HOLD still activated, press SE input until unit switches to -16dB setting
Input ~10V @ 10kHz at SE input as measured at TP9
| Enter measured level in the 10kHz Vin -16dB SE cell below |  | 10.579
| Connect AC DMM to TP7, enter measured level at 10kHz Vout in -16dB SE cell below |  | 1.675
| Switch generator to 100kHz, Enter value measured at TP9: |  | 10.577
| Connect AC DMM to TP7, adjust VC3 for DMM indication of green highlighted value |  | 1.6747

## BAL flatness
Calibration procedure Linear Audio AutoRanger v1.5 |  | Part 4 – BAL mode frequency response
Note: only do BAL flatness calibration after SE flatness calibration has been completed!
Warning: DO NOT change the contents of the green highlighted cells!
The black numbers in the J column have to be replaced with what you measure!
Step 1 -48dB cal | Place shorting BNC plug on SE input, or short TP9 to TP3
Un-select HOLD, set AR-1 to BAL mode
Set AR to HOLD then press SE input until unit switches to -48dB setting
Connect signal source to TP1 and TP3 (gnd)
Input ~10V @ 10kHz at TP1.
| Enter level measured at TP1 in 10kHz Vin -48dB BAL cell below |  | 10.573
| Connect AC DMM to TP7, enter measured level at 10kHz Vout: |  | 0.04213
| Switch generator to 100kHz, enter measured value measured at TP1: |  | 10.573
| Connect AC DMM to TP7, adjust VC2 for indication of green highlighted value: |  | 0.04213
Step 2 -32dB cal | With HOLD still activated, press SE input until unit switches to -32dB setting
Input ~10V @ 10kHz at TP1 BAL inv input as measured at TP1
| Enter measured level in the 10kHz Vin -32dB BAL cell below |  | 10.573
| Connect AC DMM to TP7, enter measured level at 10kHz Vout in -32dB BAL cell below |  | 0.2655
| Switch generator to 100kHz, enter value measured at TP1: |  | 10.571
| Connect AC DMM to TP7, adjust VC5 for DMM indication of green highlighted value |  | 0.26545
Step 2 -16dB cal | With HOLD still active, press SE input to switch to -16dB setting
Input ~10V @ 10kHz at TP1
| Enter level measured at TP1 in 10kHz Vin -16dB TP1 cell below |  | 10.569
| Connect AC DMM to TP7, enter measured level at 10kHz Vout |  | 1.674
| Switch generator to 100kHz, enter level measured at TP1 in 100kHz Vin: |  | 10.569
| Connect AC DMM to TP7, adjust VC1 for indication of green highlighted value |  | 1.674